## Slide 1: LogStat
Reporting and Logging Software for VisioStats
GKE, Original Rev 1, July 2013
First implementation: Rev 2 Aug 25, 2013
Added mods to VS-800: Rev 3: Sep 2, 2013
GK Added pictures, Rev 4: Sep 3, 2013

## Slide 2: Introduction to LogStat
1
5.3
LogStat
SQLExpert
2
8.6
N
10.
string
SupportStat
JunctionStat
VisioStat
USB cable
USB hub
Laptop
WiFi
VisioStat
VisioStat
CheckVS
LogStat is a program which,
  Connects with multiple VisioStats
  Collects and displays received information in real-time.
  Computes statistics (ave, min, max) every minute or so.
  Stores the statistics in a Database for later processing

## Slide 3: …and here it is

## Slide 4: What is included…
CheckVS
  A program which can be used to quickly setup the connections between the Laptop, Hub and VisioStats
LogStat
  A program to connect with VisioStats, display and log received information
Comms
  A utility which displays the USB/Comm ports of the laptop
TermPro or PuTTY (freeware)
  Terminal program(s) to inspect/debug incoming data
SQLExpert (free version)
  Browser of the SQLite databases.

## Slide 5: CheckVS
1. USB comm ports are listed when the program starts
2. COM14 selected
3. [Open port] pressed
4. Real-time display of raw data as received from VisioStat at port COM14
4. Real-time display of VisioStat in microAmperes (automatically adjusted)
5. Real-time display as a percent of the full scale
Start program by clicking CheckVS.py in the \LogStat directory, or
type “CheckVS” while at D:\LogStat\

## Slide 6: CheckVS connected to four VS-800s

## Slide 7: LogStat
1. Start by clicking ‘LogStat.py’ in the \LogStat directory
2. All USB ports are discovered and displayed
3. Select the USB ports where VisioStats are connected to
4. Press [Start scanning]
5. You may select the number of records to compute statistics. Must be greater than 10; default value is 60
6. Statistics counter increments fro one (1) to [Averaged Records]
7. A record is written to the database when statistics reach [Averaged Records]. [Records added] indicates the number of recors added to the database since [Start Logging ] was pressed.
Placeholder of real-time measurements (after [Start scanning] is pressed]

## Slide 8: LogStat (cont.)
8
8. Real-time data displays in micro-Amperes
9. Raw VisioStat data
10. Statistics are computed and stored into the database

## Slide 9: Database Browsing
1. Use the “SQLite Expert” to browse the database
2. Data from all USB ports are stored into the database
3. The time-stamp of data collection (and statistics) is stored
4. The average (uA) value, the minimum value (emin), and the maximum value (emax) are stored
5. The database can be
    Can be ‘read’ while LogStat ‘writes’ data
    Sorted in any fashion
    Queried using SQL
6. Other programs can be written to process the data

## Slide 10: LogStat in action…

## Slide 11: Using the System
Required programs
  Tera Term Pro (or other Terminal Program)
    Can be started from any location
    Located in the Quick Link Bar
  SQLite Expert (or other SQLite browser)
    Can be started from any location
    Located in the Quick Link Bar
  Cmd (.cmd)
    Opens a command window in the D:\LogStat directory
  Comms (.py)
    To check if USB ports have been connected
      Located in the D:\LogStat directory. Type Comms there.
  LogStat (.py)
    Starts the LogStat program
      Located in the D:\LogStat directory. Type LogStat there.
Directories
  D:\LogStat all necessary files (as above)
  D:\DropBox\LogStatDev: George’s dev directory, don’t modify

## Slide 12: Installation Hints
Wire all VS-800s first
  To the power supplies
  To the USB ports or HUB
  Turn the power on
Check connectivity using Comms
  If you do not see the expected COMM ports redo (1)
Use CheckVS to verify that each port works correctly
  If not, exit CheckVS and start TeraTerm
  Press ^C twice, and reconfigure VisioStats
Start LogStat
  Unless (1)…(3) have been completed successfully, do not attempt to run LogStat

## Slide 13: Changes to VS-800 (built 9/2/2013)

## Slide 14: VS-800 enhancements
SENSITIVITY Adjustment
  SENSITIVITY can be adjusted without pressing the SELECT pushbutton. For a given current flow, adjust until LED1+LED2 are ON and LED4 is OFF.
USB Connection
  Serial settings are 9600 baud, 8 bits, no parity. Both transmit and receive are active.
Added Serial Number Reporting
  The new format is: NNN,CCC,S,MMM
  where:
    NNN is the S/N of the device
    CCC is a wrap around record counter from 0…255
    S is the sensitivity 0…7
    MMM is the measurement from 0…255 (full scale). Use the Measurements Scale to find the current measured.
  Measurements are reported per second, independently of the sensitivity settings.

## Slide 15: VS-800 Terminal Mode
Now you can a serial TERMINAL to VS-800
  To do so, start your favorite terminal program, and while the instrument is reporting, press ^C followed by ^C within half a second.
  Examples follow

## Slide 16: Examples of Terminal Mode Commands
Basically you can modify VS-800’s configurations remotely and reboot the device.
  cmd: S
    Show EEPROM parameters
  cmd: r
    RESET. Equivalent to POWER OFF followed by POWER ON.
  cmd: q
    Quit TERMINAL mode, back to NORMAL operation.
  cmd: w 0 2
    Enable TEST mode. (Default is disabled.)
  cmd: w 0 4
    Enable LINEAR scale for the LEDs. (Default is disabled.)
  cmd: w 2 34
    Change SERIAL NUMBER to 34

## Slide 17: Summary of Modes of Operation
NORMAL MODE
  Measure current, Enable FLASHER, Enable BUZZER (if the BUZZ ON pushbutton is pressed), Transmit to the USB/UART measurements every second
CONFIGURATION MODE
  Useful to RESET to VisioStat to the default settings
  Useful to enable the TEST MODE
  While in CONFIG MODE, the VisioStat is NOT running.
  To enable the CONFIG MODE,
    Press POWER OFF
    Press and hold SELECT and press POWER ON concurrently. Release SELECT.
    Select 0 (RESET) or 1 (TEST MODE) using the SENSITIVITY
    Press POWER OFF, then POWER ON.
TERMINAL MODE
  Useful to interact with the VisioStat remotely using a terminal over USB/RS-232
  Used to modify parameters and change serial numbers
  To enable TERMINAL MODE, press ^C twice while in NORMAL MODE
    While in TERMINAL MODE, the VisioStat is NOT running.
    To quit back to NORMAL MODE, press “q” followed by RETURN
TEST MODE
  Useful to simulate current readings without connecting anywhere.
  To enable TEST MODE, use either the CONFIG MODE or TERMINAL MODE (w 0 2 command).
    You can simulate an input current from 0 to maximum full scale using the SENSITIVITY potentiometer.
    To return to NORMAL mode, use either the CONFIG MODE or TERMINAL MODE (w 0 0 command)

## Slide 18: State Diagram of Modes of Operation
NORMAL MODE
TERMINAL MODE
OFF
CONFIG MODE
TEST MODE

## Slide 19: Appendix: Related Python Modules

## Slide 20: Organization
LogStat
  LogStat.py
  CheckVS.py
  Comms
  VisioLog.sqlite
  classFrames.py
  classVisioStat.py
  classDataBase.py
  classComm1.py
  classMetter.py
PYZO
Files relating to LogStat utilities
Main program for LogStat
Utility to check VS connectivity
Command line utility to list comm ports
Database maintained by VisioStat
LogStat.py support – GUI frames
LogStat.py support – VisioStat access
LogStat.py support – Database access
CheckVS.py support – GUI frames and I/O
CheckVS.py support – Analog meter GUI
Python libraries (do not modify!)

## Slide 21: Code hierarchy
LogStat
main()
classFrames
Comms.
Measure.
Logger.
classVisioStat
VisioStat.
toeng()
classDataBase
logDB.
CheckVS
main()
classComm1
Comm1.
classMetter
Meter.
classDataBase
logDB.
Comms
main()